## Slide 1: Operation and Upgrade of the SLS RF Systems
Lukas Stingelin :: RF-Systems 1 :: Paul Scherrer Institute
26th European Synchrotron Light Source RF (ESLS-RF) workshop organized by ELETTRA November 2023

## Slide 2: SLS-Operation Statistics 2021-2022
2021:
Best beam availability ever 99.3%
Short of a 16kV transformer
Failure of reflected power measurement in SSA
Failure of 6-way splitter in SSA
2022:
Taper-foil of Undulator blocked beam-path
Failure of Magnet-PS
Incident at Super-3HC (see slides)
FPF causes vacuum-interlocks
Klystron-vacuum-PS failure
LINAC Klystron Arcing

## Slide 3: 2023 SLS Failures and Issues
Water leak at Injector LINAC Quadrupole (17h) 
Defective water-leak sensor (1h)
Issues with Super-3HC Cryoplant (17h):
  Broken valve and contactor
  Broken compressor-screw
Water leaks at Booster RF-cavity
Issue with Circulator Temperature Compensation Unit (drifting reflected power)
LINAC Broken Charging-Unit  broke power contactor and tripped fuse.
Water-leak at diagnostics beam-line

## Slide 4: Super-3HC: March 2022 Incident
Orbit problems, current drops
Vacuum increase, voltage asymmetry starts
Operator tries to increase voltages
Beam-loss (CBI)
Vacuum-controller did not interlock the beam!
Orbit off by several mm!
During machine development shift, experiments with orbit bumps.
Faillure of BPM system caused large orbit excursion
Vacuum at chamber before Super-3HC and after reached 1E-5mbar!
Asymmetry stays during accumulation and drifts over 10s of minutes to new equilibrium.

## Slide 5: Super-3HC: March 2022 Incident Reconstruction of distorted orbit and Synrad Calculation
Reconstructed orbit during the first incident:
Bending magnet Be06m is around s=137m
Super-3HC is after s=141m, its flange-to-flange length: 1.1m
After the incident: detuning of the Super-3HC changed by -5.5kHz (cav 1) and -10kHz (cav 2) for the same voltage
No obvious change in LHe consumption
Synrad: Increased power at tapers before and after the cavity
Orbid data: M. Boege, Synrad template: L. Schuz, Locations: N.Kirchgeorg
Heating due to trapped mode at valve, bellow?
Synchrotron-radition?
Quality factor degradation cannot be the explanation 

## Slide 6: Super-3HC: March 2022 Incident Investigations: Measurement of Flange Temperature and Movement
View from outer wall: beam comes from the right!
Temperature change 0…400mA:
ΔT1=+4.3°C ΔT2=+3.2°C ΔT3=+0.4°C ΔT4=+0.3°C
View from inner wall: beam comes from the left!
Longitudinal position of the flange changes with same time-constant as the voltage-drift  (after accumulation 0…400mA)
Driving offset-bumps with up to 0.5mm horizontally and 1.6mm vertically did not show any significant effect.
Installation by D. Kunz

## Slide 7: Post Mortem Pictures during SLS Disassembly
Bellow in front of the Super-3HC: 
Inspection shows a molten, welded and bent RF finger
Absorber in front of the Super-3HC: 
Some traces at vacuum chamber and absorber-block
Needs closer inspection.

## Slide 8: Super-3HC – Design of Inspection Tool
Flange upstream
Flange downstream
PMMA pipe Da=12mm / Di=10mm
Linear ball bearings
Slide: R. Rotundo 8.6.2023
 Setup for eventual inspection with endoscope
 Trolley for clean-room.

## Slide 9: LINAC: Optimized e-gun Pulser Board
 Optimized board:
Marx puls generator configuration
350V, 60ps rise-time, jitter: 15ps
Fwhm: 150ps
Courtesy: Marcos Gaspar.

## Slide 10: Last Beam of SLS1 and Ceremonial Beam-Dump

## Slide 11: SLS RF works
Done:
Modification of 500MHz Teststand for RI-cavity.
LINAC Klystron: HV-Pulse recording
LINAC Klystron HV-Pulses reduced from 6.1us to 4.8us
LINAC e-gun: pulser board upgrade to shorter bunches
In Progress:
Storage-Ring: LLRF upgrade
SLS2 RF Upgrade  See next slides

## Slide 12: More radiated X-ray power for users Less electricity consumption SLS  SLS2.0 Ee- 2.4 GeV  2.7 GeV PSR 310 kW  365 kW Welec/y 24 GWh  17 GWh Welec-WPV/y 17 GWh  15.5 GWh Key savings: Electromagnets  permanent magnets Klystrons  solid state amplifiers Standard pumps  regulated pumps for cooling Tar paper roof  PV cladded roof
≈ -30%
Power Economy SLS2.0 vs. SLS incl. PV roof
≈ -35%
PV: 1.0-1.5 GWh/y
Courtesy: SLS2 Project Management

## Slide 13: Status Storage Ring RF-Systems and Interfaces
Cavities ordered at RI (4+1 spare). 1 received and tested.
Amplifiers ordered at Cryoelectra (4x150kW)
Supports construction in progress
Concept of tunnel modification, passages and waveguide shielding finished.
Waveguides ordered (some old will be cut and re-welded)
Design of cavity vacuum components finished
Modified Teststand
RF «clean-room» operational for installation and inspection.

## Slide 14: Cavity «Accessories»
Power coupler measurement adapter
Cavity upper support frame and patch panels
Trolley and bead-pull setup
Coax-to waveguide adapter (similar to DESY-design)
Cooling distribution: built prototype
Teststand cooling

## Slide 15: First Site Acceptance Test of SLS2 Cavity
Fine-tuning of the coax-waveguide transition
Installation in the RF clean-room
Bead-Pull measurement shows asymmetry on operation-mode (as expected.)
Bake-out with hot water and electric tapes.

## Slide 16: First Site Acceptance Test of SLS2 Cavity Setup of Tuner-control:
Implemented resolution-threshold on frequency error
Fast speed for large corrections
Slow speed for small corrections

## Slide 17: First Site Acceptance Test of SLS2 Cavity
Cavity installed in the high power teststand.
Power conditioning took about 1 month!

## Slide 18: First Site Acceptance Test of SLS2 Cavity
Set of temperatures at 600kVcw, (~140 l/min)  Temperature at WG1 next to coupler reaches 58°C
Optical effects at coupler

## Slide 19: SLS2 500 MHz Phase Ref. – RF Signal Distribution
Fig. 1: Conceptual block diagram
Coax 1/2 inch,
some are isolated
Old 50 W amp. modif. With new PS to remove many spurs
SMA100B + Ultra-low phase noise option, DCFM disabled
Slide: R. Kalt, group 8417, 25.10.2023
Freq. Remote ctrl. via LAN (slow cmd rate max. 5-10 Hz) or USB (fast cmd rate max. 100-200 Hz)
Remarks:
Ideally the frequency would be ramped continously, no freq. steps to avoid the BPMs respond with artificial steps on their orbit reading (on the bw limited channels).
Want to avoid DCFM in order to get the best phase noise that is provided by the generator.
SMA100B’s analog freq. ramp option produce phase jumps at start/stop  cannot use!
Compromise to use USB with cmd rate of e.g. 100 Hz and 0.1 Hz/step  ramping rate 10 Hz/sec.
Fig 2: Phase noise of single stand alone generator SMA100B (black traces) and SLS-1 IFR2042 (green trace), the IFR2042 with DCFM enabled was not measured.

## Slide 20: LLRF 500 MHz
Status
Since March 2023 low power (50 W) test place with Elettra cavity
April-September 2023 high power testing with existing Elettra + RI cavity RF conditioning
On-going Fw/Sw/Ctrls-Integration, Layers:
  FPGA impl. on-going: DAQ, Actuator, fast RF FB-loops
  Generic IOC on-going: DAQ with synchr. time-stamps
  HLA IOC on-going: Ramping & state ctrl., resonance tuning, physical units calibrations
Fig. 1: Panels screenshot on 2023-04-25 at SLS test stand run open-loop, only cavity tuning loop active.
Slide: R. Kalt, group 8417, 25.10.2023
57 kW
566 kVp
Outlook
Implementation of Fw/Sw/IOC functionalities as defined in the requirements.
Hw series production/procurement, verification, characterization
Legend: Done, not yet done

## Slide 21: RF Interlock
Purpose: Equipment protection for an RF station as a whole. All personal safety relevant signal are directly processed at the HPRF amplifier.
Completely separated from LLRF in order to simplify the long-term maintenance.
EtherCAT is used as real-time bus communication between the devices and for control system integration (Fig. 1), ecmc as EtherCAT master / IOC.
Logic located in small Spartan-type FPGA inside the «RF ILK Unit».
Reaction time from 1 us (fast inputs, e.g. RF peak detectors) up to 2 ms (slow inputs, e.g. standard 24 V or 4-20 mA inputs).
Provides a fast sum hw-signal «RF-READY» (ok/nok = 5V/0V) towards LLRF in order to synchronize and trigger LLRF actions like: RF start, RF stop + post-mortem DAQ’s.
RF ILK
LLRF
RF-READY
OK/NOK
Slide: R. Kalt, group 8417, 25.10.2023

## Slide 22: Super-3HC Upgrade of Tuning System
+56mm
-56mm
After 20 years operation, we plan to renew the tuning-system:
Existing and spare-system has excessive torque and might self-destroy.
Upgrade and modification proposal: New parts are in colour:
- Phytron Steppermotor with planetary gear and revolver encoder
- Rollvis leadscrew
Slide: R. Rotundo 8.6.2023

## Slide 23: Super-3HC Upgrade of Tuning System
Completely extended: fits into the existing shape
Slide: R. Rotundo 8.6.2023

## Slide 24: Wir schaffen Wissen – heute für morgen
Thanks for your attention!
Questions?